OP8 & OP8 Front Toric

GP Lens for the Correction of Keratoconus

A specially designed lens with a Hi-Aspheric Base Curve for the correction of all types of keratoconus cases such as Nipple, Oval and Globe.

The lens is manufactured utilizing the state of the art CNC diamond turning lathe cutting machines to ensure accuracy, fine edge and surfaces that promise high performance in comfort and vision.

| Lens Specifications: |  |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | Material | DK | Tint |
|  |  | HDS | 58 x 10 -11 @ 35º | Blue |
|  |  | BXO | 100 x 10 -11 @ 35º | Blue, Violet |
|  |  | Optimum EX | 100 x 10 -11 @ 35º | Blue, Grey |
|  |  | Menicon Z | 186 x 10 -11 @ 35º | Blue |
| Special Treatment: |  | Plasma |  |  |
| Base Curve : |  | 5.80 to 7.60mm |  |  |
| Diameter: |  | 9.30mm (8.50 to 10.2mm are available) |  |  |
| Central Thickness: |  | 0.14 mm @ -3.00DS |  |  |
| Sphere Power: |  | Any |  |  |
| Front cyl. Power: |  | -0.50 to -3.00D at any axis |  |  |
| Back Design: |  | HI Asphericity + Blended Bevel |  |  |
| Front Design – Spheric: |  | Sphere + Lenticular |  |  |
| - Toric: |  | Toric + Lenticular + Prism Ballast |  |  |
|  |  | Marked with a dot at 270° |  |  |

Fitting Instructions:

OP8:  The fitting of this lens should be empiric.

First choice trial lens can be calculated according to the following formulas:

B.C. = K. Flat - (K. Flat - K. Steep) /3

Check fitting according to the fluorescien pattern, 3 points touch, 1.0 to 2.0mm movement, acceptable centration. Change B.C. if necessary.

Perform over - refraction, if cylinder is needed, order the OP8 FT.

OP8 FT (Front Toric):  This lens should be ordered when residual astigmatism is found after fitting the OP8 lens. This design incorporates prism ballast on the front surface for stabilizing the lens.

Back Vertex correction must be made in order to calculate the final prescription.

Lens Care:

Due to the Plasma treatment it is recommended to clean the lenses daily with a Multi Purpose System for the first 6 months, and then consult with your contact lens fitter for a more aggressive system.